## Sheet1
1. Measure reference diode with no data correction applied and running in 497 DC or AC mode, Si: 300-1100nm in 5nm steps, Ge: 800-1800nm in 5nm steps.
2. Export measurement results to Excel and paste in columns B &C (Si), and columns H & I (Ge).
3. Paste calibration data (from bcf file) of reference Si diode in column D, and reference Ge diode in column J
Wavelength (nm) | Signal through 497 (nA) | Detector responsivity (A.W-1) | Beam Power (µW) | Wavelength (nm) | Signal through 497 (nA) | Detector responsivity (A.W-1) | Beam Power (µW)
300 | 7370.2997 | 0.13031 | 56.5586 | 800 | 12423.6062 | 0.15289 | 81.2579
305 | 9062.1901 | 0.13548 | 66.8904 | 805 | 12525.2366 | 0.15797 | 79.2902
310 | 10537.7221 | 0.13904 | 75.7888 | 810 | 12638.9921 | 0.16306 | 77.5094
315 | 11834.9099 | 0.14195 | 83.3741 | 815 | 12774.9041 | 0.16822 | 75.9436
320 | 13094.2822 | 0.14433 | 90.7254 | 820 | 12950.1894 | 0.17334 | 74.7085
325 | 14657.9576 | 0.14643 | 100.1042 | 825 | 13161.1466 | 0.17838 | 73.7818
330 | 16432.7192 | 0.14797 | 111.0578 | 830 | 13423.1031 | 0.18342 | 73.1807
335 | 18516.852 | 0.14846 | 124.7237 | 835 | 13736.4224 | 0.18859 | 72.8377
340 | 20802.0732 | 0.14802 | 140.5322 | 840 | 14106.4778 | 0.19387 | 72.7607
345 | 23150.3651 | 0.14677 | 157.7325 | 845 | 14542.3546 | 0.19925 | 72.9838
350 | 25436.7113 | 0.14521 | 175.172 | 850 | 15042.2901 | 0.20472 | 73.4788
355 | 27515.8167 | 0.144 | 191.0767 | 855 | 15594.0607 | 0.21025 | 74.1677
360 | 29340.8632 | 0.14379 | 204.0497 | 860 | 16220.1583 | 0.21587 | 75.1375
365 | 31279.7189 | 0.14506 | 215.6334 | 865 | 16895.2599 | 0.22156 | 76.2576
370 | 33823.4007 | 0.14773 | 228.954 | 870 | 17629.6741 | 0.22718 | 77.6015
375 | 37093.3145 | 0.15229 | 243.5756 | 875 | 18414.1785 | 0.23268 | 79.1399
380 | 41403.0105 | 0.1579 | 262.2127 | 880 | 19233.2491 | 0.23825 | 80.7285
385 | 45704.2247 | 0.16379 | 279.0501 | 885 | 20098.2794 | 0.24407 | 82.3475
390 | 50286.603 | 0.16988 | 296.0171 | 890 | 20984.9849 | 0.24998 | 83.9463
395 | 55878.0432 | 0.17592 | 317.6378 | 895 | 21885.8644 | 0.2558 | 85.5585
400 | 51786.077 | 0.18124 | 285.7259 | 900 | 22766.5707 | 0.26156 | 87.0408
405 | 54777.9739 | 0.18627 | 294.0708 | 905 | 23689.6813 | 0.26735 | 88.6108
410 | 59814.1432 | 0.19073 | 313.6123 | 910 | 24588.5566 | 0.27316 | 90.0152
415 | 65295.6486 | 0.19477 | 335.2462 | 915 | 25489.7863 | 0.27901 | 91.3581
420 | 71142.2265 | 0.19871 | 358.0196 | 920 | 26376.6736 | 0.28494 | 92.5689
425 | 74058.3062 | 0.20252 | 365.6756 | 925 | 27233.4635 | 0.29099 | 93.5891
430 | 76130.7716 | 0.20631 | 369.0072 | 930 | 28007.9454 | 0.2971 | 94.2705
435 | 76029.8908 | 0.21023 | 361.6552 | 935 | 28660.7504 | 0.30321 | 94.5229
440 | 77095.741 | 0.21404 | 360.1979 | 940 | 29505.4585 | 0.30932 | 95.3875
445 | 76825.9466 | 0.21775 | 352.8246 | 945 | 30183.8636 | 0.31544 | 95.6889
450 | 87308.0969 | 0.22142 | 394.3061 | 950 | 30950.8324 | 0.3216 | 96.239
455 | 93443.3699 | 0.22501 | 415.2865 | 955 | 31709.7694 | 0.32784 | 96.7226
460 | 105629.1819 | 0.22849 | 462.3006 | 960 | 32432.8363 | 0.33405 | 97.0889
465 | 123408.556 | 0.23185 | 532.2852 | 965 | 33102.6793 | 0.34015 | 97.3177
470 | 139846.3249 | 0.23509 | 594.8549 | 970 | 33731.7199 | 0.34622 | 97.4274
475 | 128105.2828 | 0.23836 | 537.4499 | 975 | 34306.7199 | 0.35238 | 97.3575
480 | 120457.3631 | 0.24166 | 498.457 | 980 | 34855.1601 | 0.35862 | 97.1933
485 | 125002.3365 | 0.24509 | 510.0292 | 985 | 35361.8264 | 0.3649 | 96.9074
490 | 120963.1443 | 0.24835 | 487.0715 | 990 | 35840.4219 | 0.37116 | 96.5625
495 | 125415.5159 | 0.25138 | 498.9052 | 995 | 36283.8089 | 0.37734 | 96.1556
500 | 117877.7695 | 0.25433 | 463.4856 | 1000 | 36696.0883 | 0.38353 | 95.6788
505 | 118525.5527 | 0.25745 | 460.3773 | 1005 | 37075.4272 | 0.38981 | 95.1116
510 | 119096.3507 | 0.2606 | 457.0151 | 1010 | 37420.3205 | 0.39609 | 94.4745
515 | 120877.6951 | 0.26348 | 458.7667 | 1015 | 37742.1916 | 0.40229 | 93.819
520 | 122007.4177 | 0.26633 | 458.11 | 1020 | 38033.9652 | 0.40848 | 93.1105
525 | 122956.9435 | 0.26943 | 456.359 | 1025 | 38293.5852 | 0.41476 | 92.327
530 | 123876.7624 | 0.27254 | 454.5345 | 1030 | 38532.0872 | 0.42109 | 91.5056
535 | 124642.7536 | 0.27546 | 452.4859 | 1035 | 38741.0313 | 0.42742 | 90.639
540 | 125453.0907 | 0.27841 | 450.602 | 1040 | 38928.616 | 0.43376 | 89.7479
545 | 126863.718 | 0.28147 | 450.7236 | 1045 | 39094.9637 | 0.44011 | 88.8307
550 | 127470.6364 | 0.28439 | 448.2219 | 1050 | 39229.8192 | 0.4465 | 87.8603
555 | 128205.5378 | 0.28711 | 446.5406 | 1055 | 39341.5302 | 0.45295 | 86.8571
560 | 128447.2704 | 0.28986 | 443.14 | 1060 | 39429.915 | 0.45935 | 85.8392
565 | 128599.596 | 0.29296 | 438.9662 | 1065 | 39502.3227 | 0.46563 | 84.8364
570 | 129302.4302 | 0.29614 | 436.6273 | 1070 | 39544.5496 | 0.47188 | 83.802
575 | 129618.7401 | 0.2991 | 433.3608 | 1075 | 39577.9848 | 0.47819 | 82.767
580 | 129966.855 | 0.30194 | 430.4324 | 1080 | 39586.553 | 0.48447 | 81.7102
585 | 130394.125 | 0.30482 | 427.7692 | 1085 | 39578.6494 | 0.49068 | 80.6607
590 | 129997.8971 | 0.30763 | 422.5777 | 1090 | 39554.8612 | 0.4969 | 79.6037
595 | 128324.151 | 0.31032 | 413.5216 | 1095 | 39519.0924 | 0.50323 | 78.5303
600 | 126289.7968 | 0.31283 | 403.6949 | 1100 | 39460.3401 | 0.50971 | 77.4169
605 | 125188.3984 | 0.31618 | 395.9451 | 1105 | 212344.2143 | 0.51632 | 411.2657
610 | 124954.1283 | 0.31919 | 391.4677 | 1110 | 225787.9376 | 0.523 | 431.7136
615 | 125867.2476 | 0.32196 | 390.9346 | 1115 | 238885.2984 | 0.52972 | 450.9686
620 | 126499.1522 | 0.3246 | 389.7123 | 1120 | 251395.7918 | 0.53644 | 468.6399
625 | 125135.3741 | 0.32736 | 382.2542 | 1125 | 263603.6575 | 0.54316 | 485.3182
630 | 122740.1972 | 0.33031 | 371.5929 | 1130 | 275629.7886 | 0.54986 | 501.2745
635 | 120276.8326 | 0.33333 | 360.8359 | 1135 | 287591.3978 | 0.55653 | 516.7562
640 | 117281.3416 | 0.3363 | 348.7378 | 1140 | 299725.3239 | 0.5632 | 532.1837
645 | 116461.0863 | 0.33908 | 343.4614 | 1145 | 310717.3443 | 0.56987 | 545.2381
650 | 118008.9474 | 0.3418 | 345.2622 | 1150 | 320635.9446 | 0.57653 | 556.1508
655 | 112474.2985 | 0.34457 | 326.4211 | 1155 | 330990.7317 | 0.58313 | 567.6141
660 | 109970.1166 | 0.34741 | 316.541 | 1160 | 340892.8514 | 0.5897 | 578.0794
665 | 109863.9488 | 0.35026 | 313.6673 | 1165 | 349666.5359 | 0.59626 | 586.4295
670 | 109823.3461 | 0.35316 | 310.9743 | 1170 | 357280.8504 | 0.60275 | 592.7493
675 | 107517.2424 | 0.35586 | 302.1357 | 1175 | 364533.3052 | 0.60908 | 598.5023
680 | 107565.2838 | 0.35835 | 300.1674 | 1180 | 371122.5986 | 0.61515 | 603.3042
685 | 117766.2849 | 0.36128 | 325.9716 | 1185 | 376160.0554 | 0.62093 | 605.8022
690 | 119445.4432 | 0.36425 | 327.9257 | 1190 | 380129.0989 | 0.62652 | 606.731
695 | 112907.8388 | 0.36709 | 307.5773 | 1195 | 383577.7342 | 0.63204 | 606.8845
700 | 39499.8014 | 0.36971 | 106.841 | 1200 | 386764.7946 | 0.63752 | 606.6708
705 | 40263.4561 | 0.37244 | 108.1067 | 1205 | 389577.0609 | 0.64296 | 605.9147
710 | 40740.4363 | 0.3752 | 108.5845 | 1210 | 392270.565 | 0.64843 | 604.9581
715 | 40998.9089 | 0.378 | 108.4639 | 1215 | 394843.6379 | 0.65397 | 603.7676
720 | 41069.7788 | 0.38087 | 107.8302 | 1220 | 397119.5221 | 0.65943 | 602.2191
725 | 41020.9984 | 0.38367 | 106.9162 | 1225 | 399162.5905 | 0.66466 | 600.5525
730 | 40863.1891 | 0.38648 | 105.7316 | 1230 | 400925.4575 | 0.66971 | 598.6571
735 | 40605.2291 | 0.38933 | 104.2949 | 1235 | 402230.1137 | 0.67466 | 596.1956
740 | 40263.8584 | 0.39213 | 102.6787 | 1240 | 403117.9547 | 0.67958 | 593.1895
745 | 39856.3534 | 0.39483 | 100.9466 | 1245 | 404072.1357 | 0.68449 | 590.3247
750 | 39385.5006 | 0.39748 | 99.087 | 1250 | 405391.0971 | 0.68938 | 588.05
755 | 38843.2741 | 0.40005 | 97.0971 | 1255 | 406946.063 | 0.69422 | 586.195
760 | 38219.5204 | 0.40281 | 94.8815 | 1260 | 408520.6091 | 0.69898 | 584.455
765 | 37616.7178 | 0.40589 | 92.6769 | 1265 | 409808.4271 | 0.70366 | 582.3996
770 | 37020.5015 | 0.40867 | 90.5874 | 1270 | 410769.105 | 0.70825 | 579.9743
775 | 36370.3012 | 0.41106 | 88.4793 | 1275 | 411639.8096 | 0.7128 | 577.4938
780 | 35716.0211 | 0.41373 | 86.3274 | 1280 | 412157.4163 | 0.71743 | 574.4947
785 | 35038.9093 | 0.4164 | 84.1471 | 1285 | 412652.9992 | 0.72222 | 571.3673
790 | 34367.8594 | 0.41918 | 81.9885 | 1290 | 412892.6694 | 0.72711 | 567.8545
795 | 33738.7294 | 0.42208 | 79.9353 | 1295 | 413052.0523 | 0.73198 | 564.2964
800 | 33153.3998 | 0.425 | 78.0072 | 1300 | 412979.0962 | 0.73671 | 560.5699
805 | 32597.6223 | 0.42798 | 76.1668 | 1305 | 412815.8092 | 0.74124 | 556.9275
810 | 32131.6481 | 0.43087 | 74.5744 | 1310 | 412511.0805 | 0.74558 | 553.2762
815 | 31739.1276 | 0.43362 | 73.195 | 1315 | 412016.511 | 0.74981 | 549.4946
820 | 31447.3569 | 0.4362 | 72.0939 | 1320 | 411367.5952 | 0.75409 | 545.5174
825 | 31267.8277 | 0.4387 | 71.2743 | 1325 | 410530.0605 | 0.75854 | 541.2104
830 | 31218.5049 | 0.44147 | 70.7146 | 1330 | 409360.8558 | 0.76312 | 536.4326
835 | 31298.8639 | 0.44476 | 70.373 | 1335 | 408029.6755 | 0.76773 | 531.4783
840 | 31507.9719 | 0.44782 | 70.3585 | 1340 | 406095.4154 | 0.77231 | 525.8158
845 | 31847.8763 | 0.45004 | 70.7665 | 1345 | 401766.8366 | 0.77686 | 517.1701
850 | 32322.5856 | 0.45202 | 71.5069 | 1350 | 391989.9166 | 0.78138 | 501.6636
855 | 32905.9392 | 0.45452 | 72.3972 | 1355 | 375626.266 | 0.78592 | 477.9428
860 | 33607.7809 | 0.4573 | 73.4917 | 1360 | 360350.2512 | 0.79049 | 455.8562
865 | 34395.1255 | 0.46005 | 74.7638 | 1365 | 353497.2966 | 0.79503 | 444.6323
870 | 35264.7513 | 0.46273 | 76.2106 | 1370 | 355538.4278 | 0.7993 | 444.8111
875 | 36207.1604 | 0.46538 | 77.8012 | 1375 | 350264.281 | 0.80317 | 436.1003
880 | 37181.9079 | 0.46807 | 79.4374 | 1380 | 327040.8511 | 0.80721 | 405.1507
885 | 38203.7252 | 0.47074 | 81.156 | 1385 | 316649.4966 | 0.81185 | 390.0323
890 | 39232.0037 | 0.47337 | 82.8785 | 1390 | 327500.9692 | 0.81644 | 401.1349
895 | 40241.0328 | 0.47584 | 84.5676 | 1395 | 335340.4403 | 0.8203 | 408.8032
900 | 41216.8622 | 0.47825 | 86.1822 | 1400 | 343163.8181 | 0.824 | 416.4599
905 | 42198.1782 | 0.48063 | 87.7975 | 1405 | 352401.8526 | 0.82821 | 425.497
910 | 43123.2363 | 0.48311 | 89.2611 | 1410 | 360701.71 | 0.83279 | 433.126
915 | 44011.5273 | 0.48582 | 90.5929 | 1415 | 367794.8415 | 0.83744 | 439.191
920 | 44842.5114 | 0.48864 | 91.7692 | 1420 | 372071.0874 | 0.84209 | 441.8419
925 | 45600.2057 | 0.49141 | 92.7955 | 1425 | 374667.4657 | 0.84674 | 442.4812
930 | 46191.1291 | 0.49386 | 93.5313 | 1430 | 375295.9967 | 0.85143 | 440.7847
935 | 46565.634 | 0.49587 | 93.9061 | 1435 | 374816.3581 | 0.85618 | 437.7785
940 | 47198.9661 | 0.49777 | 94.8207 | 1440 | 374329.1199 | 0.86097 | 434.776
945 | 47550.8571 | 0.49985 | 95.1299 | 1445 | 373407.0063 | 0.86576 | 431.3031
950 | 48001.3251 | 0.5017 | 95.6772 | 1450 | 372930.8546 | 0.87053 | 428.3932
955 | 48361.4057 | 0.50286 | 96.1728 | 1455 | 372839.123 | 0.87523 | 425.9891
960 | 48626.2172 | 0.50356 | 96.5645 | 1460 | 371785.5513 | 0.87973 | 422.6123
965 | 48729.9532 | 0.50394 | 96.6989 | 1465 | 369319.5581 | 0.88395 | 417.8049
970 | 48691.988 | 0.5035 | 96.7063 | 1470 | 366872.6683 | 0.88806 | 413.1179
975 | 48478.8448 | 0.50184 | 96.6029 | 1475 | 365248.1735 | 0.89228 | 409.3445
980 | 48121.0381 | 0.49896 | 96.443 | 1480 | 364070.6837 | 0.89682 | 405.956
985 | 47596.3026 | 0.49495 | 96.1644 | 1485 | 362753.8085 | 0.90184 | 402.2363
990 | 46869.0455 | 0.48949 | 95.751 | 1490 | 361479.2824 | 0.90719 | 398.4617
995 | 45943.6387 | 0.48214 | 95.2909 | 1495 | 359822.9885 | 0.91263 | 394.2707
1000 | 44824.6777 | 0.47296 | 94.7756 | 1500 | 358076.6022 | 0.91791 | 390.0986
1005 | 43460.7416 | 0.46209 | 94.0531 | 1505 | 356219.3215 | 0.92286 | 385.9957
1010 | 41894.5998 | 0.4492 | 93.2647 | 1510 | 354226.619 | 0.92762 | 381.8669
1015 | 40085.7478 | 0.43391 | 92.3835 | 1515 | 352245.6884 | 0.93244 | 377.7689
1020 | 38057.5776 | 0.41606 | 91.4709 | 1520 | 350122.1538 | 0.93759 | 373.4266
1025 | 35828.7036 | 0.39564 | 90.5579 | 1525 | 348018.527 | 0.94323 | 368.9633
1030 | 33401.522 | 0.37273 | 89.6142 | 1530 | 345663.0409 | 0.94895 | 364.2565
1035 | 30820.6379 | 0.3479 | 88.5893 | 1535 | 343223.1545 | 0.95429 | 359.663
1040 | 28128.9607 | 0.32136 | 87.5312 | 1540 | 340387.1953 | 0.95908 | 354.9116
1045 | 25380.3343 | 0.29287 | 86.6595 | 1545 | 337246.1796 | 0.96318 | 350.1397
1050 | 22697.6812 | 0.26376 | 86.0544 | 1550 | 333737.9396 | 0.96628 | 345.3836
1055 | 20087.3062 | 0.23572 | 85.2156 | 1555 | 329286.5157 | 0.96763 | 340.3027
1060 | 17417.6797 | 0.21008 | 82.9083 | 1560 | 324173.7783 | 0.9648 | 336.0007
1065 | 15239.9302 | 0.18792 | 81.098 | 1565 | 317450.3744 | 0.95531 | 332.301
1070 | 13616.3279 | 0.16926 | 80.4475 | 1570 | 309390.4555 | 0.938 | 329.8388
1075 | 12128.789 | 0.15282 | 79.367 | 1575 | 299674.7792 | 0.91259 | 328.3784
1080 | 10780.6988 | 0.13804 | 78.0961 | 1580 | 287665.8738 | 0.88084 | 326.5817
1085 | 9561.1386 | 0.12436 | 76.8811 | 1585 | 274494.2605 | 0.84498 | 324.8526
1090 | 8470.6247 | 0.11174 | 75.8042 | 1590 | 260466.0988 | 0.80699 | 322.7629
1095 | 7478.345 | 0.1002 | 74.6357 | 1595 | 245634.1237 | 0.76876 | 319.5207
1100 | 6551.8099 | 0.089013 | 73.6054 | 1600 | 231643.0658 | 0.73213 | 316.3965
|  |  |  | 1605 | 218314.2006 | 0.69864 | 312.4865
|  |  |  | 1610 | 206084.1322 | 0.66864 | 308.2125
|  |  |  | 1615 | 195287.5257 | 0.64213 | 304.1237
|  |  |  | 1620 | 185323.9983 | 0.61872 | 299.5295
|  |  |  | 1625 | 176230.1624 | 0.59792 | 294.74
|  |  |  | 1630 | 168265.0894 | 0.57924 | 290.4954
|  |  |  | 1635 | 160671.1149 | 0.56211 | 285.8353
|  |  |  | 1640 | 153486.7585 | 0.54576 | 281.237
|  |  |  | 1645 | 146962.7768 | 0.52951 | 277.5473
|  |  |  | 1650 | 140800.6847 | 0.51342 | 274.2413
|  |  |  | 1655 | 134769.6483 | 0.4977 | 270.7826
|  |  |  | 1660 | 129111.29 | 0.48243 | 267.6264
|  |  |  | 1665 | 123792.4099 | 0.46762 | 264.7276
|  |  |  | 1670 | 118566.6025 | 0.45325 | 261.5926
|  |  |  | 1675 | 113475.3004 | 0.43926 | 258.3301
|  |  |  | 1680 | 108667.0831 | 0.4256 | 255.3298
|  |  |  | 1685 | 104054.8608 | 0.41216 | 252.4623
|  |  |  | 1690 | 99530.3541 | 0.39888 | 249.5221
|  |  |  | 1695 | 95060.8477 | 0.38574 | 246.4347
|  |  |  | 1700 | 90861.9016 | 0.37293 | 243.6446
|  |  |  | 1705 | 86818.4716 | 0.3606 | 240.7622
|  |  |  | 1710 | 82942.6408 | 0.34862 | 237.9143
|  |  |  | 1715 | 78910.917 | 0.33681 | 234.2896
|  |  |  | 1720 | 75060.5106 | 0.325 | 230.9561
|  |  |  | 1725 | 71363.6577 | 0.31306 | 227.9557
|  |  |  | 1730 | 67804.4558 | 0.3009 | 225.3373
|  |  |  | 1735 | 64331.1501 | 0.28849 | 222.9923
|  |  |  | 1740 | 60788.3513 | 0.27595 | 220.2868
|  |  |  | 1745 | 57428.4434 | 0.26343 | 218.0043
|  |  |  | 1750 | 54175.1504 | 0.25098 | 215.8588
|  |  |  | 1755 | 50968.8973 | 0.23861 | 213.61
|  |  |  | 1760 | 47851.9917 | 0.22628 | 211.4734
|  |  |  | 1765 | 44746.7357 | 0.21397 | 209.1285
|  |  |  | 1770 | 41720.3963 | 0.20181 | 206.7352
|  |  |  | 1775 | 38722.1187 | 0.18993 | 203.8754
|  |  |  | 1780 | 35804.4922 | 0.17834 | 200.7699
|  |  |  | 1785 | 32971.1229 | 0.16699 | 197.4466
|  |  |  | 1790 | 30246.1565 | 0.15587 | 194.0498
|  |  |  | 1795 | 27609.9563 | 0.14494 | 190.4924
|  |  |  | 1800 | 25072.1574 | 0.13407 | 187.0135